{"text_format":"2","workbook":"xlsx","n_sheets":1,"sheets":["清洁度需求及原型确认"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"清洁度需求及原型确认","index":0,"n_rows":130,"n_cols":9,"rows_kept":130,"cut":false,"hidden":false,"merged":[[37,2,37,7],[1,4,1,7],[2,2,2,7],[3,2,3,7],[35,4,35,7],[36,2,36,7],[104,4,104,7],[105,2,105,7],[106,2,106,7],[59,4,59,7],[60,2,60,7],[61,2,61,7],[76,4,76,7],[77,2,77,7],[78,2,78,7]]}
[]
[null,"时间","2021年10月27日 13:30 ~ 17:00","记录者"," 沙永刚"]
[null,"主题","关于清洁度需求及原型的沟通"]
[null,"人员","王新贺、陈睿、沙永刚、徐艳伟、沈成才、路婧"]
[null,"No","提出内容","提出者","确认者","确认结果","2021/12/22 会议确认内容（石峰老师）","状态"]
[null,1,"赛拓原型设计及思维导图的作成，参照的视频有两个，该视频会提供给OPTON","赛拓","欧波同","OK",null,"关闭","已处理"]
[null,2,"原型作成参照了行业标准VDA19、ISO16232","赛拓","欧波同","OK","根据客户需求考虑ISO4406、ISO4407","关闭","已处理"]
[null,3,"该系统只对接一个相机、光源、显微镜，载物台支持厦门载物台（王守壮经理提供的载物台也支持）","赛拓","欧波同","2021/11/20 林斌\n王守壮经理提供的载物台也要支持（颗粒高度无法测量）",null,"关闭","评估工作量"]
[null,4,"目前开发使用Scope5金相显微镜，而非体视显微镜，是否OK？","赛拓","欧波同","2021/11/20 林斌\n需要软件系统可同时厦门扫描台和体视显微镜？","5倍物镜视场大，速度快，","关闭","什么意思 ？"]
[null,5,"软件系统支持明场/偏光模块自动切换，但目前没有到位可自动切换的明场/偏光切换模块，现阶段开发使用手动切换模块，是否OK？","赛拓","欧波同","OK","需要自动化","关闭"]
[null,6,"软件系统支持智能光源控制控制器，但目前该控制器未到场，目前开发使用手动调整光源是否OK？","赛拓","欧波同","OK","最好自动；如果手动调整，提示用户调试是否OK","关闭"]
[null,7,"软件系统支持放大倍率识别，但目前没有倍率识别模块，因此现阶段切换物镜需要手动在软件中切换倍率，是否OK？","赛拓","欧波同","OK","不选标尺拼图有问题；建议只提供一个镜头，因此标尺固定","关闭"]
[null,8,"关于相机是基于其SDK开发的，相机SDK提供的相关设定功能都会在我们的软件中可以设定","赛拓","欧波同","OK",null,"关闭"]
[null,9,"关于载物台的软件控制，可以实现XYZ三轴联动，但目前厦门载物台XY移动时有轻微抖动","赛拓","赛拓","2021/11/20 林斌\n本硬件系统进行开发",null,"关闭"]
[null,10,"关于载物台，目前软件需求只针对单样品滤膜，不支持多工位滤膜","赛拓","欧波同","2021/11/20 林斌\n仅针对单工位扫描台开发","意义不是很大，滤膜背景不一致问题，除非自动对焦，自动白平衡","关闭"]
[null,11,"关于拍照模式，软件系统会支持黑白墨色和彩色模式","赛拓","欧波同","OK","颗粒识别灰度，对感兴趣颗粒拍照为彩色模式","关闭"]
[null,12,"软件系统内置常用标准（VDA19、ISO16232、ISO4406、ISO4407、NAS1638、SAE AS4059、DIN51455），同时支持分级修改保存","赛拓","欧波同","OK","建议只保留VDA19、ISO16232（最新版）","关闭"]
[null,13,"支持在线测量，即定位颗粒预览和拍照，需要注意：此时拍照相片只会影响报告中照片，不会替换分析结果中图库照片","赛拓","欧波同","OK","保留","关闭"]
[null,14,"载物台需要有卡位，这样不需要每次测量都去重新定位，如果没有卡位，软件支持滤膜手动定位","赛拓","欧波同","OK","系统初始大概定位，人工制定检测范围，如果自动识别检测范围最好（三点确定圆）","关闭"]
[null,15,"放大倍数100倍的情况下，样品尺寸47毫米*47毫米，显微镜视场100微米，单次拍照小图片分辨率2048*2048，偏光/明场两遍扫描（不进行对焦），至自动分析出结果，在5~6分钟内完成","赛拓","欧波同","2021/11/20 林斌\n50倍，偏光/明场两遍扫描，5分钟内；\n奥林巴斯最快100倍5分钟，蔡司最快50倍4分钟内；","50X，5分钟OK，100X，10分钟\n使用蔡司标样，每8小时校样，误差在允许范围内，在一定时期，如半年内有误差时需要做25次校样","关闭","需要调研"]
[null,16,"关于林彬与赛拓定好的原型，继续由赛拓修改或设计；","赛拓","欧波同","王新贺负责沟通确认"]
[null,17,"关于UI风格，是基于赛拓目前原型进行部分修改进行软件开发，还是完全重新设计UI风格制作原型？","赛拓","欧波同","王新贺负责沟通确认"]
[null,18,"需要赛拓两周一次提供最新源代码给欧波，由欧波同进行代码Review，赛拓进行对应修改","欧波同","赛拓","王新贺负责沟通确认",null,null,"不可以，项目验收前不提供任何源码"]
[null,19,"关于权限控制，是否需要考虑区分管理员权限与普通操作者权限（管理员可以进行如规程设置，系统硬件设置等相关高级但不常用的设置）","欧波同","欧波同","2021/11/20 林斌\n根据权限，分别控制","比较有必要的","关闭","需要调整原型及整个逻辑"]
[null,null,"规程设置时，有多个设置按钮，对于设置过的，需要有状态标识为已设置","欧波同","赛拓",null,null,null,"需要调整原型及逻辑"]
[null,null,"规程不做另存和加载(需要有导出)，创建规程时有三个选项：空白(有默认值)、加载(以往导出的)、复制(系统中现有的规程)","欧波同","赛拓",null,null,null,"需要调整原型及逻辑"]
[null,20,"原型中缺少关于硬件设置的相关画面及说明，需要赛拓补充（原型修改继续由赛拓完成的场合）","欧波同","赛拓",null,null,null,"需要调整原型及逻辑"]
[null,21,"软件系统支持两台相同型号的相机，（在软件中选择使用哪台，并非同时使用）","赛拓","欧波同","OK",null,"关闭"]
[null,22,"目前原型设计，在规程中进行的相机设置，可以保存为一个配置条目，并被其他所有可见并可以使用，是否有必要将一个规程设置的相机配置，保存为配置条目并被其他规程可见并可用？","欧波同","欧波同","2021/11/20 林斌\n相机设置各规程间可见共享；相机设置可恢复默认设置",null,"关闭","需要调整原型及逻辑"]
[null,23,"软件系统支持自动聚焦补偿（厦门载物台可以，但王守壮经理提供的扫描台没有绝对坐标，因此不能进行自动聚焦补偿(赛拓再调查)）","赛拓","欧波同","OK","一种实时聚焦，一种聚焦补偿，推荐后者对每个视场进行补偿\n前提滤膜平整","关闭"]
[null,24,"全局硬件设置中，是否有必要考虑设定一个Z轴补偿，用于对机械系统的误差补偿","欧波同","双方","不需要",null,"关闭"]
[null,25,"软件系统支持偏光/明场模块的自动切换，当没有自动切换模块，或者只有明场拍照时，系统会提示让用户切换或确认当前模块是否正确","赛拓","欧波同","OK",null,"关闭"]
[null,26,"规程设置-执行流程中的图像处理可选项是固定不可扩展的（扩展需要修改软件），前期与林斌已确认","赛拓","欧波同","2021/11/20 林斌\n目前原型中保留测量区域设置、灰度化、白平衡、阴影矫正，其他图像处理不需要（后续根据实际可能需要添加扩展），同时添加一个明场/偏光灰度差设定处理",null,"关闭","需调整原型"]
[null,27,"规程设置-执行流程，测量区域设置中，用于选定范围的线宽固定为1个像素，选定范围外区域用暗阴影表示区分","欧波同","赛拓",null,"OK",null,"需调整原型及逻辑"]
[]
[null,"时间","2021年10月28日 9:00 - 12:00","记录者"," 沙永刚"]
[null,"主题","关于清洁度需求及原型的沟通"]
[null,"人员","王新贺、陈睿、沙永刚、徐艳伟、路婧"]
[null,"No","提出内容","提出者","确认者","确认结果","2021/12/22 会议确认内容（石峰老师）","状态"]
[null,28,"原型中，参数设置-特征参数中的<宽高比>，其中高指的长度，需要将<高宽比>修改为<长宽比>","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,29,"原型中，参数设置-特征参数的可选项目是固定不可扩展的，前期与林斌已确认","赛拓","欧波同","2021/11/20 林斌\n需要追加灰度值相关特征参数(灰度计数、灰度值标准偏差、灰度值总和、灰度值平方和、最大灰度值、最小灰度值、平均灰度值)，其他特征参数后期可能会少量扩展",null,"关闭","需调整原型及逻辑"]
[null,30,"原型中，参数设置-因为测量参数添加时就完成保存，因此右下角的保存不在需要","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,null,"详细参数：组和颗粒级别名称需要命名，命名规范，可参考蔡司","欧波同","赛拓",null,null,null,"需调整原型及逻辑，并请提供具体命名"]
[null,null,"公式参数：面积、周长之类，需要增加，并且需要确定一个统一的名称","欧波同","赛拓",null,null,null,"需调整原型及逻辑，并请提供具体命名"]
[null,null,"详细参数：选择标准，位置放在最上，体现从属关系。","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,null,"详细参数：项目编号，默认自动生成，生成规则可以在主菜单的<基础设置>中设置","欧波同","双方","2021/11/20 \n可能需要包含零部件设别码、检测日期、检测方等简洁或缩写码\n根据标准整理初版（沙）",null,null,"需调整原型及逻辑，并请提供具体内容"]
[null,31,"原型中，参数设置-此处新增的测量参数条目，目前是全局规程共享的，是否有必要共享？","欧波同","欧波同","2021/11/20 林斌\n各规程全局共享",null,"关闭"]
[null,32,"关于测试报告，系统所能测试分析得到的数据，均可通过设置报告模板输出到报告中","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,33,"目前原型中，没有样品基本信息模板的设定，如样品名称、单位、制样人、制样方法等设定，设定后，会在执行检测时提示用户按照该模板进行输入，模板的设定具体可参照蔡司的<规程设置>-<输入表格>设计","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,34,"基本设置主菜单中，需要有默认用户全局信息设置，具体需要设置哪些信息和在哪使用这些信息需要确认","欧波同","双方","2021/11/20\n根据标准整理初版（沙）",null,null,"需调整原型及逻辑，并请提供具体内容"]
[null,35,"<规程设置>-<基础设置>中，关于最小颗粒尺寸和最大颗粒尺寸的设置，目前使用的单位是像素，为什么用像素？欧波同建议使用颗粒物理尺寸微米进行设定","欧波同","赛拓",null,null,null,"需调整原型及逻辑，并确认是否使用微米"]
[null,36,"<规程设置>-<基础设置>中，可以设置结果列表的显示字段，目前是参数设置中所有使用和未使用的特征参数都会显示在这里，并可选择，是否考虑只有参数设置中使用了的特征参数，才可以在这里进行选择？","欧波同","双方","2021/11/20 林斌\n在结果显示中，各字段都是可选的，在没有明显性能影响的情况下，可在分析时对全部特征参数进行提前计算(包括未使用特征参数)，这样在结果列表中选择的时候，不需要再计算",null,"关闭","需调整原型及逻辑"]
[null,37,"当同一零部件，颗粒密度较大时，需要分级过滤，分为多个滤膜进行扫描分析，由于软件系统目前不支持多滤膜同时扫描分析，因此需要软件有项目合并功能，即将多个滤膜的扫描结果进行合并统计分析和报告，目前软件系统不支持项目合并，原型需要加上该设计，实现方式可参照蔡司。","欧波同","赛拓",null,"必要的",null,"需调整原型及逻辑"]
[null,38,"执行流程时，目前设计是可以临时修改当前规程再执行的（临时修改不会影响原保存的规程），在执行流程界面，是否有必要可以临时修改规程？","欧波同","双方","2021/11/20 林斌\n执行界面不可修改规程（赛拓是否有其他考虑）","OK","关闭","需调整原型及逻辑"]
[null,39,"进入流程执行界面，需要有一个默认选中的规程（默认或锁定），目前原型没有，需要加上","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,40,"扫描分析后，需要对扫描结果进行存档处理，以便日后查看、出报告和统计分析等处理。目前原型中没有，需要加上","欧波同","赛拓",null,"OK，数据存储分两部分，1.数据，2.图片（正常扫描的图片无需保存）",null,"需调整原型及逻辑"]
[null,41,"目前软件设计，对于大图的拼接，仅仅采用了边缘模糊处理然后直接拼接，没有进行特征识别或者其他有助于准确拼接的算法，但需要软件拼图无明显拼图痕迹(接缝)，拼图无错位，不影响颗粒识别","欧波同","赛拓",null,"OK，保证一个颗粒再四个视场中时，也要能被正确识别\n可采用用以上方式矫正",null,"需调整原型及逻辑"]
[null,42,"软件系统支持对感兴趣颗粒的高度测量（手动控制软件测量）","赛拓","欧波同","OK","高度属于进阶测量，只测量感兴趣颗粒；手动测量高度OK","关闭","需调整原型及逻辑"]
[]
[null,"时间","2021年10月28日 13:30 - 15:30","记录者"," 沙永刚"]
[null,"主题","关于清洁度需求及原型的沟通"]
[null,"人员","王新贺、陈睿、沙永刚、徐艳伟、路婧"]
[null,"No","提出内容","提出者","确认者","确认结果","2021/12/22 会议确认内容（石峰老师）","状态"]
[null,43,"软件系统支持界面及及报告的国际化，目前支持中文和英文（切换语言需要重新启动系统）","赛拓","欧波同","2021/11/20 林斌\n需要同时支持中英德三种语言，软件易维护和扩展来支持其他语言","看客户群体","关闭"]
[null,44,"分析结果展示，需要提供根据分级名称(包括上下线)进行筛选的功能，方便用户对特定级别颗粒的查看和编辑。目前原型中没有，需要加上","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,45,"关于颗粒编辑，原型是用鼠标画线的方式进行连接和分割，需要修改为用鼠标选定需要连接或者分割的矩形范围，使用算法进行拟合连接或者分割。","欧波同","赛拓",null,null,null,"需调整原型及逻辑"]
[null,46,"颗粒的分类判定方法需要明确，特别是当一个颗粒中既有黑色又有金属光泽，此时如何判定金属非金属（面积占比？连接宽度？...）","赛拓","欧波同","2021/11/20 林斌\n参见Sheet<附页 - 清洁度颗粒识别确认>","参见附页红体字","关闭","需调整原型及逻辑"]
[null,47,"分析结果展示的图库界面左下的<放大比例>，指的是当选择一个颗粒并移动到待编辑颗粒时使用的放大倍数，此处的<放大比例>容易引起歧义（误以为图库图片展示比例大小），需要将此设定移动到编辑颗粒界面内。","欧波同","赛拓","2021/11/20 林斌\n缩放比例需要移动到编辑界面",null,"关闭","需调整原型及逻辑"]
[null,48,"分析结果展示的图库界面右侧的<类型>下拉框，下拉框的选择变化，直接会修改该颗粒的类型，此操作方式不直观并容易误操作，需要修改：一种方式是在下拉框右侧增加类似<确认>这样的按钮，另一种方式是将改变类型处理移动到编辑颗粒画面内，并使用我们推荐使用后一种方式处理，再一种方式为在颗粒上上点击右键进行颗粒类型变更（不能选择为自身类型）","欧波同","赛拓","2021/11/20 林斌\n类型变更：加一个确认或者应用按钮（供赛拓参考）",null,"关闭","需调整原型及逻辑"]
[null,49,"该系统主界面已经很简洁，是否还需要<基础设置>中的模块管理、工具栏设定功能？","欧波同","欧波同","2021/11/20 林斌\n不需要模块设定及工具栏设定",null,"关闭","需调整原型及逻辑"]
[null,50,"<基础设置>中的测量线设置中，是否有必要？测量线在哪会使用？会使用哪些？","欧波同","欧波同","2021/11/20 林斌\n只需要设定长度和平行线两种线型，样式只需保留字体、字号、字色、线宽、线色",null,"关闭","需调整原型及逻辑"]
[null,51," ","欧波同","赛拓"]
[null,52,"关于标尺的应用，应该在规程设置时选择并关联上现有标尺中的一个，同时在对某个颗粒观察拍照时也可以选择一个标尺（如果有倍率识别模块，则需要自动关联标尺），而不是在<基础设置>中激活和应用标尺","欧波同","赛拓",null,null,null,"需要提供全自动硬件"]
[null,53,"不同倍率拍照时，需要能够自动识别对应倍数的手动标尺","欧波同","赛拓",null,null,null,"需要提供全自动硬件"]
[null,54,"以后蔡司清洁度原图和蔡司导出图，发送给欧波同","赛拓","赛拓",null,null,null,"应该已发送"]
[]
[null,"时间","2021年10月26前","整理者"," 沙永刚"]
[null,"主题","清洁度需求理解确认（林斌前期提出的需求）"]
[null,"人员"]
[null,"No","提出内容","提出者","确认者","确认结果","2021/12/22 会议确认内容（石峰老师）","状态"]
[null,55,"<硬件对接>-<自动显微镜/手动显微镜>：可以自动识别标尺，或者可以自动识别手动添加的标尺（目镜切换，软件中标尺自动对应切换）","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"需要全自动显微镜"]
[null,56,"<硬件对接>-<相机>：拍照，曝光时间，等基本测量功能需要存在","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"拍照曝光时间在，但基本量测功能要干啥用？"]
[null,57,"<硬件对接>-<模块控制>：软件可以控制模块的自动切换，或者手动切换偏光明场的时候，软件也适配","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"需要自动控制的控制器及对应硬件"]
[null,58,"<设置功能>-<硬件参数设置>：可以对当前硬件组合（放大倍数，载物台移动步长，曝光时间，相机模式，功能模块，自动聚焦）进行保存，并在测试开始之前能一键切换到设置好的状态","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,59,"<设置功能>-<可以通过实时的预览界面来调整拼图的区域大小>：可以进行多点聚焦，用户可以设置不同区域，或者软件内置几个标准的自动聚焦区域","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,60,"<设置功能>-<拼图区域设置>：可以通过实时的预览界面来调整拼图的区域大小","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,61,"<设置功能>-<分级设置>：添加标准，以及判定，可以另存多个分级选项，并且进行调用，非选定颗粒的计算（caculate without）","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,62,"<设置功能>-<报告模板设置>：照片测量数据以及柱形图连接可以自行更改","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"需要确认，书签的模式是否合适，如果不合适，是否统一第三方报表编辑器"]
[null,63,"<设置功能>-<测试程序设置>：可以导入导出，以及删除，可以新建或者修改，或者锁定","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,64,"<设置功能>-<图像来源设置>：可以实时拼图或者离线分析图片","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,65,"<分析程序>-<结果展示>：滤膜覆盖率","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"什么是滤膜覆盖率"]
[null,66,"<分析程序>-<结果展示>：颗粒总数，不同级别颗粒数量，允许颗粒数量，判定展示，分析面积","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的\n※分析结果中可根据参数字段，如面积等进行筛选"]
[null,67,"<分析程序>-<结果展示>：清洁度代码展示","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的\n※根据归一化参数选择生成清洁度代码：数量、面积、体积"]
[null,68,"<分析程序>-<结果展示>：按照不同的特征参数可以重新排列颗粒","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"是否为图库-排序功能"]
[null,69,"<分析程序>-<结果展示>：报告图片能够设置是否显示标尺","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"是否为图库-显示标尺功能"]
[null,70,"<分析程序>-<分析执行程序>：交互式执行，可以显示或隐藏部分固定的设置","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,71,"<分析程序>-<分析执行程序>：可以设置标样单次扫描程序","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的\n※对表样单次扫描，以检测软硬件系统是否有问题",null,null,"不明白什么意思"]
[null,72,"<分析程序>-<报告输出>：可以显示颗粒真实颜色，不光是黑白颗粒","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的",null,null,"是否黑白颗粒，取决于拍摄时使用的相机参数"]
[null,73,"<其他功能>-<工作流>：可以自己编写工作流","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的，需要赛拓再确认","可以设置为隐藏或者不隐藏（熟练者可能不需要看到）"]
[null,74,"<其他功能>-<颗粒高度>：颗粒高度可以选择感兴趣颗粒测量（最高点，最低点）自动测量最好","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的（支持高度手动测量，是否支持自动测量需要再讨论）",null,null,"需要确认"]
[null,75,"<其他功能>-<数量换算>：按照不同的计算方式来进行颗粒数量的换算（零件数，表面积，体积）","林斌","赛拓","2021/11/20 林斌\n左侧内容是必要的"]
[null,76,"<其他功能>-<中英文切换>：最好是可以进行中英文的切换","林斌","赛拓","2021/11/20 林斌\n需用同时支持中英德三种语言，对后续其他语言支持，软件易于扩展"]
[null,77,"<其他功能>-<多操作者使用快捷方式>：多个快捷方式，每个快捷方式里的参数独立设置","林斌","赛拓","2021/11/20 林斌\n不需要","建议操作人员不能修改任何参数，让操作人可调的参数越少越好"]
[]
[null,"时间","2021年10月26前","整理者"," 沙永刚"]
[null,"主题","清洁度性能及效率确认（赛拓原型设计中提出）"]
[null,"人员"]
[null,"No","提出内容","提出者","确认者","确认结果","2021/12/22 会议确认内容（石峰老师）","状态"]
[null,78,"处理图片速度","赛拓","林斌","①扫描拍照速度：50X，3分钟\n②颗粒识别速度：50X，1分钟\n③颗粒测量速度：50X，1分钟\n总体原则：从扫描开始，到自动分析出报告，50X不超过5分钟即可",null,"关闭","需要调研"]
[null,79,"处理图片效果","赛拓","林斌","①大图拼接效果：不错位、无阴影拼缝\n②图库小图效果：颗粒清晰、无马赛克、缩放比例一致（同类颗粒）\n③报告图片效果：可选彩色、颗粒清晰、无马赛克",null,"关闭","需要调研"]
[null,80,"处理的准确性","赛拓","林斌","①颗粒分类准确性：人工观察可明显识别的颗粒类型，软件必须正确\n②颗粒分析准确性：50μm以上，准确率≥99.5%；5μm~50μm，准确率≥95%",null,"关闭","需要调研"]
[null,81,"测试时的图片","赛拓","林斌","①图片类型：赛拓确认\n②图片大小：赛拓确认",null,null,"需要调研"]
[null,82,"采用拼接算法","赛拓","林斌","①接缝：不能有明显拼接印痕\n②错误：不能错位或重叠\n③分析：不影响颗粒识别与测量",null,"关闭","需要调研"]
[null,83,"拼图速度要求","赛拓","林斌","①分辨率：2464*2056（V16 相机像素分辨率）\n②拼图数：50倍下19行23列大概400张图，两次扫描800张图片，图片大小3.48G\n③速度要求：拼图速度50X下5min以内；100X下10min以内，\n                      20X下扫描时间2min、16X下6行7列、一共80张图左右；",null,null,"需要调研"]
[null,84,"处理大图能力","赛拓","林斌","①分析50μm颗粒，5μm/pixel，6*7张照片：两遍84张图\n②分析5μm颗粒，1μm/pixel，19*23张照片：大概400张图，两遍800张图",null,null,"需要调研"]
[null,85,"清洁度硬件确认","赛拓","林斌","①明场/偏光模块：滤膜夹持器以及滤膜盖板一套，王部长协调硬件\n②载物台：厦门载物台\n                 （130*85扫描台，考虑要放置两个滤膜，现在只能单一滤膜）\n③多个控制器：王部长协调硬件"]
[null,86,"同一滤膜软件重复性测试","赛拓","林斌","①颗粒分类准确性：人工观察可明显识别的颗粒类型，软件必须正确\n②颗粒分析准确性：50μm以上，准确率≥99.5%；5μm~50μm，准确率≥95%",null,"关闭","需要调研"]
[null,87,"操作性要求","赛拓","林斌","与ax对比：体验更简洁、在ax的基础上进行微调  ",null,"关闭"]
[null,88,"功能性要求","赛拓","林斌","非常用的可以不添加到系统中（未添加的重要功能需要欧波同确认）",null,"关闭"]
[null,89,"颗粒高度测量","赛拓","林斌","需要支持颗粒高度测量",null,"关闭"]
[]
[]
[null,null,"首先是准确性，保证准确性的前提下其次是速度"]
[null,null,"稳定性：标准样件，反复检测一定遍数，如25遍，测试结果稳定一致"]
[]
[null,null,"过程变化分析也很重要有意义，BI分析对管理者有意义"]
[null,null,"可视化展示通过服务提供，各个终端展示，展示不要求实时化"]
[null,null,"数据网络化，云端化，加密算法"]
[]
[null,null,"UI需要美化，操作越简洁越好，设定越少越好"]