{"text_format":"2","workbook":"xlsx","n_sheets":3,"sheets":["问题列表","Sheet2","Sheet1"],"not_rendered":{"charts":0,"images":2,"macros":0,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"问题列表","index":0,"n_rows":32,"n_cols":10,"rows_kept":32,"cut":false,"hidden":false,"merged":[[1,2,1,3],[2,2,2,3],[3,2,3,3]]}
[]
[null,"会议","清洁度会议"]
[null,"时间","2022/11/09 14:00-16:30"]
[null,"人员","孙总、王新贺、林斌、刘洪冰、宋珊珊、沙永刚、王琦"]
[null,"No","模块","内容","结论","区分","优先级","责任者","工作量(h)","状态"]
[null,1,"全体","主工作区中的窗口，在较小屏幕（如笔记本）部分项目表示不全","画面使用滚动条","问题","中","宋珊珊",12]
[null,2,"规程设置","规程设置-图片来源：倍率列表中选择项目应该系统标尺下的所有倍率","倍率选择删除，只保留标尺即可，因为标尺决定分辨率并关联了倍率","改善","中","沙永刚",0.5,"完成"]
[null,3,"规程设置","规程设置-图片来源：阴影矫正选项中不要空选项","删除空项","问题","中","沙永刚",0.5,"完成"]
[null,4,"规程设置\n项目管理","规程浏览、项目浏览加右键快捷菜单功能","需要追加","改善","中","宋珊珊",8]
[null,5,"检测","检测结果-图库模式/分级模式的Grid需要可手动改变列宽","需要修改，无特殊要求的都需要按此展开修改","改善","中","宋珊珊",1]
[null,6,"检测","提供颗粒参看功能（类似编辑画面）","提供查看模式，两种实现方式：一种是单独做个查看画面，可以缩放，囊括更多范围；另一种是在编辑画面，提供查看模式（不可编辑只可参看），编辑画面需提供缩放，以便囊括更多范围\n\n目前考虑，第二种方式","新需求","中","沙永刚",8]
[null,7,"检测","提供拼接后大图查看功能（可放大查看细节），参照av的评估图","加分项，需要增加该功能，参照av实现，难点是大图加载性能问题","新需求","低",null,null,"二期"]
[null,8,"检测","颗粒图库显示颗粒数量一多，出结果的时候卡在数据处理\n\n关于颗粒图片，如图库浏览、颗粒编辑、报告颗粒图片，都是使用时拼接大图实时切割下来的，优点是切割幅度可以随意设定和动态调整，这种方式是否可以？\n另一个种是保存所有颗粒切割图，使用时直接使用该图，缺点是：一次检测会占用更多磁盘空间（可能30万颗粒的小图），切割幅度不能随意调整（调整重新从大图切割保存），\n\n是否需要在图库模式的缩略图和颗粒信息列表中显示所有颗粒：目前看显示所有会有性能问题，如显示30万颗粒","保持现状：不保存颗粒小图，需要时从大图实时切割颗粒小图；不需要在页面显示所有颗粒；\n\n但需要进行以下优化和修改：（图库模式性能优化相关）\n1.当切割图片超过一定限值时，进行压缩切割\n2.提供分页功能：每页可设定展示较少数量图片，分页显示","改善","高","王琦",24]
[null,9,"规程设置","规程设置-拼图设置：选择聚焦补偿的聚焦点数时，需要在右侧预览区域标记聚焦点位置","加分项，需要增加该功能","新需求","低",null,null,"二期"]
[null,10,"全体","所有画面的ComboBox控件下拉框的宽度需要修改","要修改，全部画面展开检查","问题","中","沙永刚",1,"完成"]
[null,11,"规程设置","规程设置-执行流程-物相提取：\n两个阈值场合，不同阈值识别的轮廓需要用不同颜色轮廓标识","现状OK","-","-","-","-","-"]
[null,12,"系统设置","物镜转盘设定：需要可动态设置增删改查倍率（不应让客户手动修改配置文件）","要修改，可手动设定倍率，而不仅仅是从下拉框选择\n设定物镜转盘后，其他有倍率选择画面，其倍率选择项目需要从物镜转盘内容中读取","改善","低","宋珊珊",16]
[null,13,"设置","缓存存在哪个路径，C盘每分析一个项目都会变大，怎么清空缓存","现状OK。可通过点击：设置-磁盘清理，跳转到图片存储目录，可手动清空（图片缓存区位置可由用户进行设置指定）","-","-","-","-","-"]
[null,13,"报告","缓存存在哪个路径，C盘每分析一个项目都会变大，怎么清空缓存","缓存区图片被清空后，不能图库模式查看图片，不可以出报告。\n报告需要修改：可参照图库模式，当对应图片不存在是，提示图库被清除，不能出报告信息（参照点击图库模式按钮处理）","问题","高","宋珊珊",2]
[null,14,"规程设置\n检测","物相提取这里还是不能交互，无法根据实际情况调整提取阈值","要修改，提供检测时用户的交互设定阈值功能，难点是大图加载","新需求","高","沙永刚",40]
[null,15,"规程设置\n检测","物相提取增加吸管功能","加分项，需要增加该功能","新需求","低",null,null,"二期"]
[null,16,"检测","颗粒分割这里点击完成以后处理速度很慢；\n\n更改类型还是每改一个类型更新全部，重新计算功能没有起到作用；","需要修改：（编辑与重新计算相关）\n\n更改类型处理类似现状的编辑颗粒，每更改一个类型，就重新进行全部统计分析。该处理需要对应：更改后，不立即全部更新，而是重新计算后统一更新","改善","高","王琦",40]
[null,18,"检测","有些颗粒连接分割不上","现状OK。连接不上可能是一下两种情况：\n1.同时选定连接范围的颗粒数量大于或等于3个了（一次连接操作只可连接两个颗粒，超过时不进行处理）；\n2.选定需要连接的颗粒中不包含当前颗粒（编辑画面只允许编辑当前颗粒，连接的颗粒必须包含当前颗粒）；","-","-","-","-","-"]
[null,18,"检测","分割和连接，点击使用有时不起作用，需要再次点击一次按钮；有时候点击后切割不起作用，换不同位置反复测试，有时候就可以分割，有时不能分割","需要调查和修改","问题","高","刘洪冰",16]
[null,19,"报告","这个版本点生成报告没反应","新版本OK","-","-","-","-","-"]
[null,20,"规程设置","相机扫描跟本地图片规程是否考虑可以合并","需要合并，大分支节点只有用户规程和系统规程区分","改善","中","沙永刚",4,"完成"]
[null,21,"规程设置\n检测","各级分析结果通过与否是跟哪个数据比对","现状OK。根据规程设置-分级设置中指定【允许数量使用的归一化参数】一致","-","-","-","-","-"]
[null,22,"规程设置\n检测","关于归一化参数选择需要可设定","加分项，需要增加该功能","新需求","低",null,null,"二期"]
[null,23,"检测","没有某一颗粒不参与计算的选项","需要修改：只需在进行全部类型统计时，不参与计算的类型不参与全部类型的统计分析运算，但不参与全部类型计算的类型，自己的相关测量数据，及图库展示正常，与现状一致；另外全部类型的滤膜覆盖率包含不参与计算的颗粒类型","问题","高","沙永刚",8,"完成"]
[null,24,"检测","图库模式，提供Grid列头点击排序（图库要联动）","不可点击Grid列头排序","改善","中","王琦",1]
[null,25,"算法","测试压缩前后对比数据，得出数据","需要对比测试","改善","低","沙永刚",8]
[null,26,"全体","多次测量分析等操作后，内存是否释放","需要优化分析，如有问题需要进行改善","改善","中","宋珊珊",16]

{"sheet":"Sheet2","index":1,"n_rows":14,"n_cols":17,"rows_kept":14,"cut":false,"hidden":false,"merged":[]}
{"$rows_empty":7}
[{"$nulls":7},"位置","↓",null,"倍率",null,null,"名字",null,null,"添加"]
[{"$nulls":8},1]
[{"$nulls":8},2]
[{"$nulls":8},3]
[{"$nulls":8},4]
[{"$nulls":8},5]
[{"$nulls":8},6]

{"sheet":"Sheet1","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}