{"text_format":"2","workbook":"xlsx","n_sheets":1,"sheets":["4月14日会议内容"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"4月14日会议内容","index":0,"n_rows":38,"n_cols":5,"rows_kept":38,"cut":false,"hidden":false,"merged":[[1,2,1,4],[2,2,2,4],[3,2,3,4],[36,2,37,2],[16,2,19,2],[20,2,21,2],[22,2,28,2],[29,2,30,2],[31,2,35,2]]}
[]
[null,"会议","清洁度原型审查"]
[null,"时间","2022年4月14日 09::00-12:00"]
[null,"人员","王新贺、林斌、宋珊珊、沙永刚、王琦"]
[null,"序号","模块","内容","结论"]
[null,1,null,"报表最终通过哪种方式实现？\n（1）Excel\n（2）报表设计器","待确认"]
[null,2,null,"亮度百分比需要保留，考虑调到外面，不放到规程中","OK"]
[null,3,"规程↓","相机参数中分开偏光和明场两项","OK"]
[null,4,null,"拼图区域有必有增加矩形。","OK"]
[null,5,null,"聚集补偿参照：阵或是洋葱皮方式，软件自动分布点，也允许客户在自动分布基础上手动增加点","先期软件自动分配点"]
[null,6,null,"拼图设置，左侧载物台实时预览考虑去掉。","去掉"]
[null,7,null,"信息模板中允许选择图片，加到报告中。","OK"]
[null,8,"项目信息模板","部件管理中，在执行检测过程，可以在下拉列表中选择部件，支持直接添加新部件。并保存到数据库。目前先考虑选择，不可编辑。编辑在部件管理中完成。","OK"]
[null,9,"信息模板中","提供 数量、表面程、体积；三项客户都不输，按颗粒数量计算的方式","OK"]
[null,10,"执行流程","二值个部分参数李伟的软件，在这上增加空洞和碎削","OK"]
[null,11,null,"测量区域保留，（灰度化、白平衡可以去掉）","OK"]
[null,12,"测量参数","测量参数中加一下优先级，金属优先还是非金属优先","OK"]
[null,13,null,"预增加一些常用的参数。","OK"]
[null,14,null,"字段考虑加一个英文名，方便公式计算。函数这一列要去掉。","OK"]
[null,15,null,"灰度用最大灰度和平均灰度","OK"]
[null,16,"分级设置","“使用参数”取自测量参数中列出的参数","OK"]
[null,17,null,"分级表格换成Tab页，按颗粒类型","OK"]
[null,18,"检测↓","光源调整提示界面去掉","OK"]
[null,19,null,"扫描过程中不可以停或是在扫描过程中更换聚集方式。这个位置去掉聚集方式，只显示扫描进度就可以。不需要实时预览窗口","OK"]
[null,20,null,"暂停和继续去掉","OK"]
[null,21,null,"检测过程中不需要再弹出拼图设置了，因为在规律中已经进行了设置","OK"]
[null,22,null,"白平衡那个也去掉","OK"]
[null,23,null,"物相提取参数李伟的，简化","OK"]
[null,24,null,"执行时偏光和明场这个也去掉","OK"]
[null,25,"颗粒编辑","颗粒编辑页面增加一个刷新或是重新计算按钮。","OK"]
[null,26,null,"颗粒编辑中增加一下添加颗粒的功能。","OK"]
[null,27,"高度测量","去掉移动载物台，去掉X和Y值 移动的功能","OK"]
[null,28,null,"曝光时间增加一个手动曝光的功能，增加曝光时间的设置","OK"]
[null,29,null,"批量测量高度功能，后继考虑增加","先期手动，后续再考虑批量自动"]
[null,30,null,"替换照片增加一个景深叠加（最好是自动叠加）的功能，使照片更清楚。后继考虑","先期直接拍照，后续考虑景深叠加"]
[null,31,null,"真正的测量高度的操作，路径太深需要点多步。在结果图库显示界面或工具栏上增加一个“高度测量”的按钮。参照李伟的软件中高清替换与高度测量在一个界面中。","OK"]
[null,32,"设置↓","去掉网格设置","OK"]
[null,33,null,"显示设置去掉。在规程中的颗粒显示中的显示设置，初始给定一个默认值 。","OK"]